# Movement Baseline — what the literature says the assessment should measure

Research input for a clinical decision. Not clinical advice, and not a substitute for PT sign-off. Compiled 2026-10-14 to give a physical therapist something concrete to react to, rather than a blank page.

Companion to [MOVEMENT_BASELINE_PLAN.md](MOVEMENT_BASELINE_PLAN.md) (the engineering plan) and [CLINICAL_GUARDRAILS.md](CLINICAL_GUARDRAILS.md) (what our measurements can and cannot claim).

This is a different question from D1 in [HANDOFF.md](HANDOFF.md). D1 is the *canonical

joint-key list* — the literal strings (hipAngleMin vs hipFlexAngle) that two codebases

disagree on. That is an engineering task with no clinical content. This document is the

clinical question: which joints and movements a baseline assessment should contain.

## 1. The clinical gold standard already exists

The Fugl-Meyer Assessment (FMA) is the stroke-specific standard for motor impairment. It is performance-based, stroke-specific, and has five domains: motor function, sensory function, balance, joint range of motion, and joint pain.

| Motor score | 0–100 (66 upper extremity + 34 lower extremity) |
| --- | --- |
| Joint ROM domain | 0–44 points, across 8 joints |
| Scoring | 3-point ordinal per item — 0 = cannot perform, 1 = partial, 2 = full |
| Joints covered | shoulder, elbow, forearm, wrist, hand, hip, knee, ankle |

Two things follow for us:

There is no need to invent an assessment. A validated instrument exists, is widely used in

stroke trials, and already has a ROM domain. The design question is which *subset* of it a phone camera can honestly support.

The FMA is ordinal, not goniometric. It scores "cannot / partial / full", not degrees. That

is a much weaker claim than a measured angle — and notably closer to what our system can actually defend. Worth considering whether the Movement Baseline should report bands rather than degrees.

## 2. What actually goes wrong after a stroke

This determines which joints are worth assessing at all.

About 52% of people develop at least one major contracture within six months of stroke, and contractures can develop fully within four weeks.

The joints most commonly affected:

| Joint / motion | Why it matters |
| --- | --- |
| Ankle dorsiflexion | Limited DF is one of the main causes of functional limitation after stroke. Passive DF in stroke is roughly half that of healthy adults. Drives foot drop and gait |
| Shoulder external rotation | Part of the hemiplegic shoulder picture; post-stroke shoulder problems reported from 7% to 81%, with greater active shoulder motion associated with lower incidence |
| Elbow extension | Flexor synergy; a common contracture |
| Forearm supination | Common contracture |
| Wrist and finger extension | Common contracture |
| Hip internal rotation | Common contracture |

Note the collision with our hardware: three of the six most commonly contractured motions — forearm supination, wrist extension, finger extension — are not measurable by body-pose models at exercise distance. MoveNet and BlazePose give one wrist keypoint and no forearm rotation. Section 5 deals with this honestly.

## 3. What a phone camera can actually measure

This is the constraint that should shape the assessment set, and the evidence is mixed by joint.

### Where markerless pose estimation does well

MediaPipe against goniometer ground truth in healthy volunteers: R² = 0.988, MAPE 1.54% —

but in controlled conditions.

Stroke gait parameters: correlations ≥ 0.92 for most parameters, dropping to 0.40–0.66

for swing time.

### Where it does not

A 2025 validation specifically on elbow ROM in a clinical population:

| Measurement | Markerless error | Physician visual estimate |
| --- | --- | --- |
| Elbow flexion | 5.1° (HSMR) / 7.3° (RTMW) | 5.4° |
| Elbow extension | 7.8° / 6.5° | 5.2° |

The authors' conclusion: neither markerless method reached clinically acceptable accuracy for elbow extension. Flexion was roughly as good as a physician eyeballing it — which is a fair benchmark, but not goniometry.

Three failure modes they identify map exactly onto what we have measured ourselves:

Out-of-plane motion — substantial errors when participants internally rotated the upper arm.

(We measured the same thing: front-view hip angle during a seated march implied ~167° of hip flexion against an AAOS normal of 120°.)

Biomechanical model limits at end-range and hyperextension.

Laboratory conditions — uniform lighting, minimal background clutter. A patient's living

room is neither.

A broader review is blunter: *accurately quantifying angular metrics such as range of motion would require much more tracking precision than current open-source pose estimation models offer.*

### The camera setup that validation used

Worth copying, because it is a published, defensible spec:

RGB camera at 0.9 m height, facing the sagittal plane on the affected side, at 3 m

distance, 30 fps.

This independently supports the camera-pitch check added on 2026-10-14 — camera height is a real experimental variable, not a detail, and ours was previously unchecked.

## 4. The number that should govern the whole design

Minimal clinically important difference (MCID) for ROM in chronic stroke:

| Joint | MCID (affected side) |
| --- | --- |
| Knee, sagittal ROM | 8.48° |
| Hip, flexion–extension | 5.81° |
| Shoulder, ankle | not established in stroke |

Put that beside the measurement error from §3 — roughly 5–8° for a well-conditioned joint under laboratory conditions:

| Joint | MCID | Our likely error | Verdict |
| --- | --- | --- | --- |
| Knee | 8.48° | ~5–8° | Marginal — real change is detectable, but only just |
| Hip | 5.81° | ~5–8° | Below the noise floor — error is as large as the effect |
| Shoulder | unknown | large excursion, in-plane → likely best case | Plausible |
| Ankle | unknown | worst case (see §5) | Not defensible today |

This is the single most important finding in this document. For the hip, the measurement error is comparable to the clinically meaningful difference — so "your hip range improved" is not a defensible claim from a single phone camera, at any amount of software polish.

It does not mean the baseline is useless. It means:

The baseline is sound as an operational threshold (scaling each patient's rep gate), which is

what it already does well, and which tolerates a few degrees of error.

It is not sound as a clinical progress measurement for every joint. Report trend and

direction, with a change threshold set above measurement error — not a degree count.

That matches what CLINICAL_GUARDRAILS.md already requires: *present ROM as tracking and trending, never as measurement.*

## 5. The ankle problem

Ankle dorsiflexion is simultaneously the most clinically important motion in this list and the one we measure worst.

Clinically: limited DF is a main driver of post-stroke functional limitation, passive DF is about half of healthy, and plantarflexor spasticity routinely reduces it further.

Technically, from our own measurements (CLINICAL_GUARDRAILS.md):

Both detectors resize their person-ROI before inference (MoveNet 192×192, BlazePose 256×256), so

the foot occupies a small fixed fraction of model input no matter how close the patient stands.

A 0.05–0.15 footPitchNorm swing is roughly a 3–10 pixel vertical difference between the two

least reliable landmarks in the model.

footPitchNorm is a vertical image difference, so camera roll injects about 0.11 at 10° of

tilt against a real signal of 0.05–0.15.

MediaPipe defaults visibility to 1.0 when the model omits it, so these keypoints can carry

unearned confidence.

Recommendation: include ankle dorsiflexion in the assessment because it matters clinically, but record it as observed/not observed with a confidence flag, not as a degree value — and flag it to the PT as the measurement most in need of either a different method (a side-on close-up trial at a fixed distance) or acceptance that it is qualitative.

## 6. Proposed assessment set

Ranked by *clinical importance × measurability*. Every row needs PT review; the right-hand columns are what engineering can honestly support.

| # | Trial | Joint / motion | View | Our feature | Confidence | Clinical basis |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Arm raise to the side, slow | shoulder abduction | front and side | shoulderAngleL/R | High — large excursion, large joint, in-plane | Hemiplegic shoulder, 7–81% incidence |
| 2 | Elbow bend and straighten | elbow flexion / extension | side (sagittal, affected side) | elbowAngleL/R | Medium flexion, Low extension | Elbow extension is a common contracture; extension is where markerless failed |
| 3 | Seated knee extension | knee, open chain | side | kneeAngleL/R | Medium-high | MCID 8.48° is near our resolution |
| 4 | Sit to stand / partial squat | knee, closed chain | side | kneeAngleL/R, kneeAngleMin | Medium-high | Weight-bearing differs from open chain |
| 5 | Seated knee lift | hip flexion | side (never front) | hipAngleL/R | Low — MCID below our error | Hip IR/flexion contracture common |
| 6 | Toes up / heel up, seated | ankle DF/PF, non-weight-bearing | side, close | ankleAngleToe, footPitchNorm | Low — qualitative only | Most common and most functionally important |
| 7 | Toes up / heel up, standing | ankle DF/PF, weight-bearing | side, close | as above | Low — qualitative only | WB and NWB differ by more than double |
| 8 | Quiet standing | trunk / posture | front and side | trunkAngle, trunkAngleMin | Medium | Genu recurvatum affects 40–68% of post-stroke walkers |

That is 8 trials, several with two views — realistically 10–12 captures. Not the "5–6 exercises" originally imagined. Worth knowing before promising a patient a short session.

### Setup to specify, from the validation literature

camera at roughly chest height (~0.9 m), which our new checkCameraPitch now checks

sagittal view on the affected side for all the limb-flexion trials

~3 m distance

even lighting, uncluttered background — both named as limits of the lab validation

## 7. What we should not pretend to assess

Stating these explicitly is what keeps the product honest:

| Motion | Why not |
| --- | --- |
| Forearm supination / pronation | No forearm rotation in any body-pose model. A common post-stroke contracture we simply cannot see |
| Wrist extension | One wrist keypoint, no hand orientation |
| Finger extension | Not in the body-pose skeleton at all |
| Shoulder external rotation | Rotation about the limb's long axis is invisible to a 2D projection. Named as a most-common contracture |
| Hip internal rotation | Same reason |
| Spasticity / tone | Requires passive movement by a clinician (Modified Ashworth). Not observable from a patient moving themselves |

Four of the six most commonly contractured motions after stroke are in this table. The Movement Baseline therefore cannot be presented as a complete ROM assessment. It is a partial, camera-based screen of the motions a phone can see — which is still useful, and much better than calibrating from the first two reps of a treatment set, but it is not the FMA.

## 8. Questions for the PT

Framed so they can be answered quickly, with our constraint attached to each.

**Is an ordinal band (cannot / partial / full, as the FMA scores it) more defensible than a

degree value for our purposes?** Our measurement error is 5–8°, and hip MCID is 5.81°.

Is a standing weight-bearing ankle trial worth the fall risk for the population, given we can

only report it qualitatively?

Which single joint matters most if the assessment must be short? Literature points to ankle

dorsiflexion, which is our weakest measurement — we would value a second opinion.

Should the assessment capture the unaffected side as a within-patient reference? MCID data is

reported separately for affected (8.48° knee) and unaffected (6.81°) sides, implying both are routinely measured.

Re-assessment interval — 90 days proposed. Reasonable for chronic stroke?

Is it acceptable to omit forearm, wrist and hand entirely (§7), or does that make the

baseline clinically misleading?

## Sources

Fugl-Meyer Assessment — [Shirley Ryan AbilityLab](https://www.sralab.org/rehabilitation-measures/fugl-meyer-assessment-motor-recovery-after-stroke) · [StrokEngine](https://strokengine.ca/en/assessments/fugl-meyer-assessment-of-sensorimotor-recovery-after-stroke-fma/) · [Gladstone et al., critical review (SAGE)](https://journals.sagepub.com/doi/pdf/10.1177/154596802401105171) · [Fugl-Meyer Assessment of Sensorimotor Function After Stroke (Stroke/AHA)](https://www.ahajournals.org/doi/10.1161/strokeaha.110.592766)

Contracture prevalence after stroke — [Half of the adults who present to hospital with stroke develop at least one contracture within six months (J Physiother)](https://www.sciencedirect.com/science/article/pii/S1836955312700711) · [Incidence of contractures by joint](https://www.researchgate.net/figure/Incidence-proportion-of-contractures-by-joint-as-measured-with-the-contracture-scale_tbl2_221838960) · [WHO clinical guideline on stroke rehabilitation](https://extranet.who.int/ncdccs/Data/MNG_D1_2.%20Rehabilitation%20guideline%20of%20Stroke.pdf)

Ankle dorsiflexion after stroke — [Ankle joint mobilization, dorsiflexion and gait: systematic review and meta-analysis](https://www.researchgate.net/publication/350142117_The_Effects_of_Ankle_Joint_Mobilization_on_Dorsiflexion_Range_and_Gait_Parameters_in_Chronic_Stroke_Survivors_A_Systematic_Review_and_Meta-analysis) · [Floss band, ankle ROM, balance and gait in chronic stroke](https://doi.org/10.3390/healthcare12232384)

Markerless pose estimation validity — [Validation of markerless human pose estimation for clinical assessment of elbow ROM (PLOS ONE)](https://journals.plos.org/plosone/article?id=10.1371%2Fjournal.pone.0353801) · [Reliability and validity for hip and knee ROM (PMC11783685)](https://pmc.ncbi.nlm.nih.gov/articles/PMC11783685/) · [Markerless joint angle estimation using MediaPipe (Springer)](https://link.springer.com/article/10.1007/s11042-026-21256-z) · [Assessment of monocular pose estimation models for clinical movement analysis (Scientific Reports)](https://www.nature.com/articles/s41598-025-22626-7) · [Shoulder abduction angle with posture-estimation AI (PMC10416158)](https://pmc.ncbi.nlm.nih.gov/articles/PMC10416158/)

MCID for ROM in stroke — [Estimating MCID for knee ROM after stroke (J Clin Med)](https://www.ncbi.nlm.nih.gov/pmc/articles/PMC7602397/) · [Establishing MCID for sagittal hip ROM in chronic stroke (PMC8443407)](https://www.ncbi.nlm.nih.gov/pmc/articles/PMC8443407/) · [MCID of the lower-extremity Fugl-Meyer in chronic stroke](https://www.researchgate.net/publication/301343626_Minimal_clinically_important_difference_of_the_lower-extremity_fugl-meyer_assessment_in_chronic-stroke)

Telerehabilitation assessment batteries — [Group telerehabilitation for balance and mobility after stroke (PMC12785289)](https://pmc.ncbi.nlm.nih.gov/articles/PMC12785289) · [H-GRASP upper limb home exercise program (PMC5399812)](https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5399812/) · [Telerehabilitation therapeutic exercise in chronic stroke, RCT protocol (Life)](https://doi.org/10.3390/life15121905)